S u m m a r y

# ALTERNATIVE SITE TESTING OF BLOOD GLUCOSE IN CHILDREN AND ADOLESCENTS WITH TYPE 1 DIABETES

V. Škrabić, J. Jakšić, T. Matić, Z. Kačić, M. Šitum

Several previous studies in adults with type 1 diabetes have documented that glucose levels obtained at alternative sites and  at fingertip are not the same, a difference was noticed  after a meal, after insulin application or during and  after exercise.

This study was done specifically to look at alternative site testing (AST) to determine if the site differences previously observed in adults are also found in children and adolescents. Glucose measurements were made using OneTouch Ultra System at fingertip and forearm sites, before breakfast and dinner, but at least two hours after insulin application and two hours after a meal. Twenty-four diabetics attending a summer camp in Dalmatia participated in the testing.  At the completion of the camp, the children were questioned about their experience with AST. A total of 217 forearm measurements were made. Linear regression analysis showed that forearm glucose values were highly correlated with fingertip values (p<0.001, R=0.96). Using Clarke Error Grid analysis with fingertip readings as the reference values, 89.9% of values fell in zone  A, 8.3% in zone B, 2% fell in  zone D. Parkes Error Grid analysis showed  90.3% of values in zone A, 10% in zone B and none in  zone C and D.

Although 37% of the diabetics examined had a slight subjective feeling of  pain, 46% did not feel pain using AST, 25% emphasized the simplicity of measuring blood glucose using AST and only 29% will use AST more frequently at home in regards to using it at the summer camp. 71% will periodically or rarely use AST (41% because it requires more time and 58% because it was harder to obtain blood for measurement or visible stab wounds were present).

Glucose values obtained from forearm sites in children and adolescents studied in real life at the diabetic camp correlated closely with fingertip values. This suggests that when children follow the current guidlines for AST largely developed in adults, the results are reliable and can safely be utilized.

Descriptors:  DIABETES  MELLITUS, TYPE 1 – blood;  BLOOD  GLUCOSE – analysis;  BLOOD  GLUCOSE  SELF-MONITORING – methods

Primljeno/Received: 30. 3. 2007.

Prihvaćeno/Accepted: 10. 9. 2007.